## Sheet1
[8 empty] | LOCATION AND FUNCTIONAL DESCRIPTIONS |  |  |  | EQUIPMENT DETAILS | [38 empty] | NEW OR EXISTING
[13 empty] | Physical Size |  |  | Clearances |  |  |  |  | Electrical | [8 empty] | Data & Alarm |  | Plumbing / Piping | [13 empty] | HVAC | [10 empty] | Budget
Primary Header | Secondary Header | Area Or Department | Reference Line# | COM | Floor Level | Room # | Equipment ID Tag # | Equipment Description | Cut Sheets on Onehub | Manufacturer | Model # | Mounting | Width or Length (Inches) | Depth (Inches) | Height (Inches) | Side (Inches) | Front (Inches) | Rear (Inches) | Top (Inches) | Est. Shipping Weight (Lbs.) | Voltage | Amps (load, approx.) | Phase | Dedicated Circuit | Normal Power | Optional Standby Power | Uninterruptable Power (UPS) | TVSS | NEMA config | Data | Monitoring/Alarm | Lab Air (LA) 100 psig | Lab Vacuum (V) 26 "Hg | Dental Air | Dental Vac (DV) 10" Hg | Natural Gas (NG) | Carbon Dioxide (CO2) | Nitrogen (N2) | Oxygen (O2) | Nitrous Oxide | Special Gas | Hot/Cold Water | Drain | Pure Water | Steam (100 psig) | Process Vent | Exhaust Duct Conn. (In.) | Exhaust CFM | Exhaust SP (" w.c.) | Special Exhaust | Existing to be Relocated | New | Procured by Installed by | Existing Location (Campus, Building, Rm. ##) | Existing Equipment Condition Assessment | Net to Purchase | Funding Source | FF&E Costs - Group II Major Moveable Equip. | Departmental or Operational - Group III Minor Movable Equipment | Competitive process
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Tissue Culture Lab | Dr. Gray's LVL & Vivarium | 120 | Dr. Gray | P2 | P2N021DA | INC-P2N021DA-1 & 2 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 39.4 80 | [5 empty] | 120 120 | 12 12 | 1 1 | X X | N | X X | N | Y | Std. | [7 empty] | Y Y | [13 empty] | N | Y | OF/OI | NA | NA | 4 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Tissue Culture Lab | Dr. Gray's LVL & Vivarium | 121 | Dr. Gray | P2 | P2N021DA | LRF-P2N021DA-1 | Refrigerator/Freezer | Y | Fisher Scientific Isotemp General Purpose | 11 651 120, upright, 10 cu. ft., 11651120 3764A | Floor | 31 | 31 | 60 | 3 | Open | 6 | Open | 300 | 120 | 5.8 | 1 | N | Y | N | N |  | Std. | --- | ? | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Tissue Culture Lab | Dr. Gray's LVL & Vivarium | 123 | Dr. Gray | P2 | P2N021DA | WB-P2N021DA-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 7 | [5 empty] | 120 | 4 | 1 | N | [26 empty] | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Tissue Culture Lab | Dr. Gray's LVL & Vivarium | 124 | Dr. Gray | P2 | P2N021DA | CEN-P2N021DA-1 | Microcentrifuge | Y | Eppendorf | Model #TBD, -4 C, 14,000 RPM | Bench | 14 | 10 | 17 | [5 empty] | 120 | 1050w | 1 | [27 empty] | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Ante Room | Dr. Gray's LVL & Vivarium | 128 | Dr. Gray | P2 | P2N021 |  | Autoclave, portable | N | Thermo Fisher | 02-964-9 | Special | 20 | 20 | 15 | [5 empty] | 120 | 1400w | 1 | Y | Y | [25 empty] | N | Y | OF/OI | NA | NA | 1 | FFE-4 | Group II |  | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Ante Room | Dr. Gray's LVL & Vivarium | 129 | Dr. Gray | P2 | P2N021 | LF | Freezer, -80C, with no alarm system s | Y | Thermo Scientific | UxF60086A, 120V | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | Y | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 132 | Dr. Gray | P2 | P2N021 | LR | LABH-20-RF; flammables refrigerator | Y | LabRepCo | LABH-20-RF | Floor | 32 | 32 | 71 | Min. | Open | Min. | Open |  | 120 | 5 | 1 | N | Y | N | N | N | Std. | n | ? | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 134 | Dr. Gray | P2 | P2N021 | CEN-P2N021-1 | Centrifuge, ultra | Y | Thermo Scientific Sorvall WX | WX90 (TF cat #46901) | Floor | 32 | 28 | 40 | [5 empty] | 208 | 20 | 1 | Y | Y | N | N | N | [22 empty] | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 135 | Dr. Gray | P2 | P2N021 |  | Incubator | Y | Fisher Scientific Isotemp std lab | Catalog #51028065 | Bench | 55 | 35 | 38 | [5 empty] | 120 |  | 1 | Y | Y | N | N | N | Std. | --- | Y | --- | --- | --- | --- | --- | N | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 136 | Dr. Gray | P2 | P2N021 | WPS-P2N021-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-4 | Group II |  | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 137 | Dr. Gray | P2 | P2N021 | OV | Oven, bench top | Y | Blue-M | 11-650-901, 3.75 cu. Ft. | Bench | 18 | 18 | 20 | [5 empty] | 120 | 18 | 1 | Y | Y | N | N | N | Std. | [21 empty] | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Wet Prep Lab | Dr. Gray's LVL & Vivarium | 138 | Dr. Gray | P2 | P2N021 | EMS-P2N021 | Monitoring & alarm system, lab equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Low Vibration Lab (LVL) - Dr. Gray's microscopy | OCSSB - Microtomes Lab | Dr. Gray's LVL & Vivarium | 139 | Dr. Gray | P2 | P2N021A |  | Cryostat | Y | TBD | 3050S | Floor | 32 | 32 | 36 | [5 empty] | 208 | 1800 VA | 1 | Y | Y | N | N | Y | ?? | [21 empty] | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Vivarium | GENERAL NOTE: For budgeting and accuracy, pricing for portable equipment is carried in this GENERAL section. Equipment is also shown for each room for physical coordination purposes. | Group III | 161 | CLSB | P2 | NA | BT | Truck, heavy duty bulk, autoclaveable | N | Alternative Design Mfgr. (Janette) | BT-60x32x68, S/S, with covers | Floor | 60 | 32 | 68 | --- | --- | --- | --- |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 8 | OHSU |  | Group III | Lab
Vivarium | GENERAL NOTE: For budgeting and accuracy, pricing for portable equipment is carried in this GENERAL section. Equipment is also shown for each room for physical coordination purposes. | Group III | 167 | CLSB | P2 | NA |  | Vapor decontamination system, portable | N | Bioguell | Clarus C, Hydrogen Peroxide |  |  |  |  | --- | --- | --- | --- |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Vivarium | OCSSB - Conv. Imaging | Inter-Institute Vivarium Phase I | 227 |  | P2 | P2N058 |  | IVIS Quantum FX (low dose x-ray tomography) | Y | Caliper Life Sciences | IVIS Quantum FX | Floor | 39 | 37 | 58 | Min. | Open | Min. | Open | 300 | 120 |  | 1 | Y | [26 empty] | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 278 | CLSB | 1 | 1N046 | INC-1N046-1 INC-1N046-2 INC-1N046-3 | Isotemp Benchtop Incubators, 3.75 CF (no CO2, must have 2 shelves per incubator) | Y | Thermo Fisher | Catalog #TBD | Bench | 24 | 26 | 33 | Min. | Open | Min. | Open |  | 120 | 2.7 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 279 | CLSB | 1 | 1N046 | CEN-1N046-1 | Centrifuge, ultra | Y | Eppendorf | Model #TBD, 14,000 RPM | Bench | 15.2 | 25.6 | 11.6 | Min. | Open | Min. | Open |  | 230 | 1050 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 280 | CLSB | 1 | 1N046 | CEN-1N046-1 thru 6 | Microcentrifuge | Y | Thermo Fisher Accuspin | 75002431 | Bench | 10 | 9 | 14 | Min. | Open | Min. | Open |  | 120 | 200 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 281 | CLSB | 1 | 1N046 | WB-1N046-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | Min. | Open | Min. | Open |  | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 282 | CLSB | 1 | 1N046 | PCR-1N046-1 PCR-1N046-2 | PCR Thermal Cycler | Y | Bio-Rad MyCycler | 170-9703 | Bench | 16.4 | 10.4 | 10.8 | Min. | Open | Min. | Open |  | 120 | 670 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 284 | CLSB | 1 | 1N046 | DR-1N046-1 | Refrigerator, deli case | Y | Thermo Fisher | TBD, Capacity 61 CF | Floor | 71 | 30 | 79 | Min. | Open | Min. | Open |  | 120 | 12 | 1 | Y | N | Y | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 285 | CLSB | 1 | 1N046 | LF-1N046-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | TBD | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | Y | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 287 | CLSB | 1 | 1N046 |  | Gel Electrophoresis, horizontal, wide mini-sub cell GT system | Y | Bio-Rad | 1704485EDU | Bench | 25 | 13 | 7 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 288 | CLSB | 1 | 1N046 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell | Y | Bio-Rad | 165-8000EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 289 | CLSB | 1 | 1N046 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell casting module | Y | Bio-Rad | 165-8010EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 290 | CLSB | 1 | 1N046 | EPS-1N046-1 through 6 | Electrophoresis Power Supply | Y | Bio-Rad | 164-5050 | Bench | 11 | 12 | 6 | Open | Open | Open | Open |  | 120 | 2 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 291 | CLSB | 1 | 1N046 | EB410-1N046-1 &2 | Balance, electronic, 410 grams x 0.01 gram accuracy, student grade | Y | Ohous Adventure Pro-Precision | 01-921-19 AV412-US | Bench | 9 | 12 | 12 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 292 | CLSB | 1 | 1N046 | LS-1N046-1 | Shaker, lab, standard non-refrigerated, digital | Y | Thermo Scientific Max Q 4000 orbital | SHKE 4000 | Bench | 18 | 18 |  | Open | Open | Open | Open | 167 | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 293 | CLSB | 1 | 1N046 | HP | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Isotemp | 11-200-49SH | Bench | 11 | 8 | 5 | Open | Open | Open | Open |  | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 294 | CLSB | 1 | 1N046 | UVS-1N046-1 | Spectrophotometer, Genesys 20 Visible | Y | Fisher Scientific Spectronic 20+ | S99607 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 18 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 296 | CLSB | 1 | 1N046 | WB-1N046-2 thru 7 | Water bath, 5 L | Y | Thermo Scientific | 15-460-S5Q, with SS cover | Bench | 14 | 11 | 8 | Open | Open | Open | Open |  | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 297 | CLSB | 1 | 1N046 |  | Microscope, lab (24 per lab plus 3 spares per room) | Y | Leica DM750 | 13613510 | Bench | 12 | 12 | 16 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 27 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 301 | CLSB | 1 | 1N046 |  | CAMAG UV BOX | N | Fisher | NC9576741 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | ? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 302 | CLSB | 1 | 1N046 |  | UV box, variable intensity (dual light) | N | Fisher | SDBLT88AQ900 | Bench | [38 empty] | N | Y | OF/OI | NA | NA | 5 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 303 | CLSB | 1 | 1N046 | LRO-1N046-1, 3 &3 | Rotator, lab | N | Fisher/Barnstad | 2309Q, 9" x 9" platform | Bench | 12 | 12 | 4 | Open | Open | Open | Open |  | 120 | ?? | 1 | ? | Y | N | N | Y | std | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 304 | CLSB | 1 | 1N046 | TRAN-1N046-1, thru 4 | Transilluminator, TFM-20V, 20 x 20 cm filter size | N | Fisher | UV97-0243-01 | Bench | [38 empty] | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 305 | CLSB | 1 | 1N046 |  | Jar, anaerobic | Y | Fisher Scientific | L-60608, large | Bench | 8 | 8 | 12 | Open | Open | Open | Open | Min. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N046 | PSU Micro/Cell/Recomb DNA Labs | 307 | CLSB | 1 | 1N046 |  | Balance, electronic, 121 grams x 1MG accuracy, top loading | N | Mettler-Toledo | XP1203S | Bench | [38 empty] | N | Y |  | NA | NA | 6 |  |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 308 |  | 1 | 1N044 | INC-1N044-1 INC-1N044-2 INC-1N044-3 | Isotemp Benchtop Incubators, 3.75 CF (no CO2, must have 2 shelves per incubator) | Y | Thermo Fisher | Catalog #TBD | Bench | 24 | 26 | 33 | [5 empty] | 120 | 2.7 | 1 | N | Y | N | N | Y | std. | --- | ? | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 309 |  | 1 | 1N044 | CEN-1N044-1 | Centrifuge, ultra | Y | Eppendorf | Model #TBD, 14,000 RPM | Bench | 15.2 | 25.6 | 11.6 | Min. | Open | Min. | Open |  | 230 | 1050 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 310 |  | 1 | 1N044 | CEN-1N044-1 thru 6 | Microcentrifuge | Y | Thermo Fisher Accuspin | 75002431 | Bench | 10 | 9 | 14 | Open | Open | Open | Open |  | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 311 |  | 1 | 1N044 | WB-1N044-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | Open | Open | Open | Open |  | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 312 |  | 1 | 1N044 | PCR-1N044-1 PCR-1N044-2 | PCR Thermal Cycler | Y | Bio-Rad MyCycler | 170-9703 | Bench | 16.4 | 10.4 | 10.8 | Open | Open | Open | Open |  | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 314 |  | 1 | 1N044 | DR-1N044-1 | Refrigerator, deli case | Y | Thermo Fisher | TBD, Capacity 61 CF | Floor | 71 | 30 | 79 | Open | Open | Open | Open |  | 120 | 12 | 1 | [6 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 315 |  | 1 | 1N044 | LF-1N044-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | TBD | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | N | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 317 |  | 1 | 1N044 |  | Gel Electrophoresis, horizontal, wide mini-sub cell GT system | Y | Bio-Rad | 1704485EDU | Bench | 25 | 13 | 7 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 318 |  | 1 | 1N044 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell | Y | Bio-Rad | 165-8000EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 319 |  | 1 | 1N044 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell casting module | Y | Bio-Rad | 165-8010EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 320 |  | 1 | 1N044 | EPS-1N044-1 through 6 | Electrophoresis Power Supply | Y | Bio-Rad | 164-5050 | Bench | 11 | 12 | 6 | [5 empty] | 120 | 2 | 1 | N | Y | N | N | Y | std. |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 321 |  | 1 | 1N044 | EB410-1N044-1 &2 | Balance, electronic, 410 grams x 0.01 gram accuracy, student grade | Y | Ohous Adventure Pro-Precision | EB220-1N043-1 &2 AV412-US | Bench | 9 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 322 |  | 1 | 1N044 | LS-1N044-1 | Shaker, lab, standard non-refrigerated, digital | Y | Thermo Scientific Max Q 4000 orbital | SHKE 4000 | Bench | 18 | 18 | [5 empty] | 167 | 120 | min | 1 | N | Y | N | N | Y | std. |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 323 |  | 1 | 1N044 | HP-1N044-1 thru 12 | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Isotemp | 11-200-49SH | Bench | 11 | 8 | 5 | Open | Open | Open | Open |  | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 324 |  | 1 | 1N044 | UVS-1N044-1 thru 3 | Spectronic 200 Spectrophotometer | Y | Thermo Scientific | S06240 | Bench | [8 empty] | 120 | min | 1 | N | Y | N | N | Y | std. |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 325 |  | 1 | 1N044 | UVSM-1N044-1 thru 3 | Spectrophotometer, UV mini 1240 | N | Fisher Scientific | 20-264-0092 | Bench | 15 | 17 | 11 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 326 |  | 1 | 1N044 | WB5-1N044-2 thru 7 | Water bath, 5 L | Y | Thermo Scientific | 15-460-S5Q, with SS cover | Bench | 14 | 11 | 8 | Open | Open | Open | Open |  | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 327 |  | 1 | 1N044 |  | Microscope, lab (24 per lab plus 3 spares per room) | Y | Leica DM750 | 13613510 | Bench | 12 | 12 | 16 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 27 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 328 |  | 1 | 1N044 |  | Microscope, epi-fluorescent | N | Leica DM750 | 13613511 | Bench | 12 | 12 | 16 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 330 |  | 1 | 1N044 | SV-1N044-1 | Concentrator kit, SpeedVac > ISS110 DNA | N | Thermo Scientific | 20-548-134 | Bench | [8 empty] | 120 | ??? | 1 | N | Y | N | N | Y | std. |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 331 |  | 1 | 1N044 |  | CAMAG UV BOX | N | Fisher | NC9576741 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | ? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 332 |  | 1 | 1N044 |  | UV box, variable intensity (dual light) | N | Fisher | SDBLT88AQ900 | Bench | [19 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 5 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 333 |  | 1 | 1N044 | LRO-1N044-1, 2 &3 | Rotator, lab | N | Fisher/Barnstad | 2309Q, 9" x 9" platform | Bench | 12 | 12 | 4 | Open | Open | Open | Open |  | 120 | ?? | 1 | ? | Y | N | N | Y | std | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 334 |  | 1 | 1N044 | TRAN-1N044-1, thru 4 | Transilluminator, TFM-20V, 20 x 20 cm filter size | N | Fisher | UV97-0243-01 | Bench | [19 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 335 | CLSB | 1 | 1N046 |  | Jar, anaerobic | Y | Fisher Scientific | L-60608, large | Bench | 8 | 8 | 12 | Open | Open | Open | Open | Min. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N044 | PSU Micro/Cell/Recomb DNA Labs | 337 |  | 1 | 1N044 |  | Air jacketed CO2 Incubators |  | VWR | 98000-364 | Bench | [43 empty] | 6 |  |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 338 |  | 1 | 1N056 | INC-1N056-1 INC-1N056-2 INC-1N056-3 | Isotemp Benchtop Incubators, 3.75 CF (no CO2, must have 2 shelves per incubator) | Y | Thermo Fisher | Catalog #TBD | Bench | 24 | 26 | 33 | [5 empty] | 120 | 2.7 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | N | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 339 |  | 1 | 1N056 | CEN-1N056-1 | Centrifuge, ultra | Y | Eppendorf | 5430R, 14,000 RPM | Bench | 15.2 | 25.6 | 11.6 | Min. | Open | Min. | Open |  | 230 | 1050 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 340 |  | 1 | 1N056 | CEN-1N056-1 thru 6 | Microcentrifuge | Y | Thermo Fisher Accuspin | 75002431 | Bench | 10 | 9 | 14 | [5 empty] | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 341 |  | 1 | 1N056 | WB-1N056-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | [5 empty] | 120 | ?? | 1 | N | Y | N | N | Y | std. |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 342 |  | 1 | 1N056 | PCR-1N056-1 PCR-1N056-2 | PCR Thermal Cycler | Y | Bio-Rad MyCycler | 170-9703 | Bench | 16.4 | 10.4 | 10.8 | [5 empty] | 120 | ?? | 1 | N | Y | N | N | Y | std. |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 344 |  | 1 | 1N056 | DR-1N056-1 | Refrigerator, deli case | Y | Thermo Fisher | TBD, Capacity 61 CF | Floor | 71 | 30 | 79 | [5 empty] | 120 | 12 | 1 | [27 empty] | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 345 |  | 1 | 1N056 | LF-1N056-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | MF34SS-SARE-TS | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | N | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 347 |  | 1 | 1N056 |  | Gel Electrophoresis, horizontal, wide mini-sub cell GT system | Y | Bio-Rad | 1704485EDU | Bench | 25 | 13 | 7 | [5 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 348 |  | 1 | 1N056 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell | Y | Bio-Rad | 165-8000EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open | [10 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 349 |  | 1 | 1N056 |  | Gel Electrophoresis, vertical, mini-PROTEAN tetra cell casting module | Y | Bio-Rad | 165-8010EDU | Bench | 4.8 | 6.4 | 7.2 | Open | Open | Open | Open | [10 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 350 |  | 1 | 1N056 | EPS-1N056-1 through 6 | Electrophoresis Power Supply | Y | Bio-Rad | 164-5050 | Bench | 11 | 12 | 6 | Open | Open | Open | Open |  | 120 | 2 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 351 |  | 1 | 1N056 | EB410-1N056-1 &2 | Balance, electronic, 410 grams x 0.01 gram accuracy, student grade | Y | Ohous Adventure Pro-Precision | 01-921-19 AV412-US | Bench | 9 | 12 | 12 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 352 |  | 1 | 1N056 | LS-1N056-1 | Shaker, lab, standard non-refrigerated, digital | Y | Thermo Scientific Max Q 4000 orbital | SHKE 4000 | Bench | 18 | 18 |  | Open | Open | Open | Open | 167 | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 353 |  | 1 | 1N056 | HP-1N056-1 thru 12 | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Isotemp | 11-200-49SH | Bench | 11 | 8 | 5 | Open | Open | Open | Open |  | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 354 |  | 1 | 1N056 | UVS-1N056-1 thru 3 | Spectronic 200 Spectrophotometer | Y | Thermo Scientific | S06240 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 355 |  | 1 | 1N056 | UVSM-1N056-1 thru 3 | Spectrophotometer, UV mini 1240 | N | Fisher Scientific | 20-264-0092 | Bench | 15 | 17 | 11 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 356 |  | 1 | 1N056 | WB5-1N056-2 thru 7 | Water bath, 5 L | Y | Thermo Scientific | 15-460-S5Q, with SS cover | Bench | 14 | 11 | 8 | Open | Open | Open | Open |  | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 357 |  | 1 | 1N056 |  | Microscope, lab (24 per lab plus 3 spares per room) | Y | Leica DM750 | 13613510 | Bench | 12 | 12 | 16 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 27 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 358 |  | 1 | 1N056 |  | Microscope, epi-fluorescent | N | Leica DM750 | 13613511 | Bench | 12 | 12 | 16 | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 360 |  | 1 | 1N056 | SV-1N056-1 | Concentrator kit, SpeedVac > ISS110 DNA | N | Thermo Scientific | 20-548-134 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | ??? | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 361 |  | 1 | 1N056 |  | CAMAG UV BOX | N | Fisher | NC9576741 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | ? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 362 |  | 1 | 1N056 |  | UV box, variable intensity (dual light) | N | Fisher | SDBLT88AQ900 | Bench |  |  |  | Open | Open | Open | Open | [31 empty] | N | Y | OF/OI | NA | NA | 5 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 363 |  | 1 | 1N056 | LRO-1N056-1, 2 &3 | Rotator, lab | N | Fisher/Barnstad | 2309Q, 9" x 9" platform | Bench | 12 | 12 | 4 | Open | Open | Open | Open |  | 120 | ?? | 1 | ? | Y | N | N | Y | std | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 364 |  | 1 | 1N056 | TRAN-1N056-1, thru 4 | Transilluminator, TFM-20V, 20 x 20 cm filter size | N | Fisher | UV97-0243-01 |  |  |  |  | Open | Open | Open | Open | [31 empty] | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Micro/Cell/ Recombinant DNA Teaching Labs | Room 1N056 | PSU Micro/Cell/Recomb DNA Labs | 365 | CLSB | 1 | 1N046 |  | Jar, anaerobic | Y | Fisher Scientific | L-60608, large | Bench | 8 | 8 | 12 | Open | Open | Open | Open | Min. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 371 | CLSB | 1 | 1N043 | SS-1N043-1 | Shaker, stacked, incubated, no CO2 | Y | Thermo Fisher | SHKE-8000 | Floor | 50 | 30 | 25 |  |  |  |  | 440 | 120 |  | 1 | ? | Y | N | N | N | std. | --- | --- | --- | --- | --- | --- | --- | N | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 374 | CLSB | 1 | 1N043 | LF-1N043-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | MF34SS-SARE-TS | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | N | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 376 |  | 1 | 1N043 | ULF-1N043-1 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | Y | N | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 377 |  | 1 | 1N043 | EMS | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 378 |  | 1 | 1N043 | WPS-1N043-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 379 |  | 1 | 1N043 | INC-1N043-1 | Isotemp Benchtop Incubators, 3.75 CF (no CO2, must have 2 shelves per incubator) | Y | Thermo Fisher | Catalog #TBD | Bench | 24 | 26 | 33 | Min. | Open | Min. | open |  | 120 | 2.7 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | N | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 380 |  | 1 | 1N043 | WB-1N043-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 15.48 | 15 | 8 | [5 empty] | 120 | 300 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 381 |  | 1 | 1N043 | HP-1N043-1 thru 6 | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Cimarec Basic | HP194825 | Bench | 11 | 8 | 5 | Open | Open | Open | Open |  | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 382 |  | 1 | 1N043 |  | Pipette, Pipettman Neo P200N | Y | Gilson | F144565 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 383 |  | 1 | 1N043 | P1000-1N043-1 thru 6 | Pipette, Pipettman Neo P1000N | Y | Gilson | F144566 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 384 |  | 1 | 1N043 |  | Pipette, Pipettman Neo P10N | Y | Gilson | F144562 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 385 |  | 1 | 1N043 |  | Pipette, Pipettman Neo P2N | Y | Gilson | F144561 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 386 |  | 1 | 1N043 | UPP-1N0431 & 2 | Pump, unispense peristaltic | Y | Wheaton | 13-683-25 | Bench | 12 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 387 |  | 1 | 1N043 |  | pH meter, Bench top, pH/ISE multiparameter | Y | Thermo Scientific/ Orion | 13-645-542 STARA2146 | Bench | 12 | 12 | 6 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 388 |  | 1 | 1N043 | EB121-1N043-1 & 2 | Balance, electronic, 121 grams x 1MG accuracy, top loading | Y | Mettler-Toledo | XP1203S | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU Level 1 Biology Prep Lab |  | PSU Level 1 Bio Prep Lab | 389 |  | 1 | 1N043 | EB220-1N043-1 &2 | Balance, electronic, 220 grams x 0.001 gram accuracy | Y | Mettler-Toledo | ML203E | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Anatomy & Physiology Teaching Labs (3 labs on Level 2) | Rooms 2N108A, 2N103 and 2N109 - Teaching Labs | PSU Anatomy & Physiology Labs | 392 |  | 2 | 2N108A, 2N103, 2N111 |  | Microscope, lab (24 per lab) | Y | Leica DM500 with accessories | cat #13612310 | Bench | [8 empty] | 120 | min | 1 | N | Y | N | N | N | std. | [21 empty] | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Chemistry Teaching Labs (5 labs on Level 2) | Rooms 2N056, 2N066, 2N089, 2N099 and 2N111 | PSU General Chemistry Teaching Labs | 400 |  | 2 | 2N056, 2N066, 2N089, 2N099, 2N111 | EB220-2N056-1 -3 EB220-2N066-1 -3 EB220-2N089-1 -3 EB220-2N099-1 -3 EB220-2N111-1 -3 | Balance, electronic, 220 grams x 0.001 gram accuracy | Y | Mettler-Toledo | ML203E | Bench | 9 | 12 | 12 | M | Open | M | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 15 | PSU |  | Group III | Lab
PSU General Chemistry Teaching Labs (5 labs on Level 2) | Rooms 2N056, 2N066, 2N089, 2N099 and 2N111 | PSU General Chemistry Teaching Labs | 401 |  | 2 | 2N056, 2N066, 2N089, 2N099, 2N111 | BTO-2N056-1 BTO-2N066-1 BTO-2N089-1 BTO-2N099-1 BTO-2N111-1 | Oven, bench top | Y | Blue-M/Lindberg | 510 288 72, 5.0 CF (one per lab) | Bench | 26 | 24 | 40 | M | Open | M | Open | 145 | 120 | 1800 watts | 1 | Y | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 5 | PSU |  | Group III | Lab
PSU Level 2 Chemistry Solutions Lab (1) |  | PSU Solutions Lab & Stockroom | 414 |  | 2 | 2N057 | WPS-2N057-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
PSU Level 2 Chemistry Solutions Lab (1) |  | PSU Solutions Lab & Stockroom | 415 |  | 2 | 2N057 | EB121-2N057-1 | Balance, electronic, 121 grams x 1MG accuracy, top loading | Y | Mettler-Toledo | XP1203S | Bench | 9 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Chemistry Solutions Lab (1) |  | PSU Solutions Lab & Stockroom | 416 |  | 2 | 2N057 | EB210-1N057-1 | Balance, electronic, 210 grams x 0.01 gram accuracy | Y | Mettler-Toledo | XP204S | Bench | [8 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
|  | PSU Solutions Lab & Stockroom | 418 |  | 2 | 2N057 | CB40-2N057-1-6 | Carboy, 40 L | Y | VWR | 89170-780 HDPE | Bench | [8 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 10 | PSU |  | Group III | Lab
PSU Advanced Chemistry Lab (1) |  | PSU Advanced Chemistry Lab | 426 | CLSB | 2 | 2N079 | BTO-2N079-1 | Oven, bench top | Y | Blue-M/Lindberg | 510 288 72, 5.0 CF (one per lab) | Bench | 26 | 24 | 40 | M | Open | M | Open | 145 | 120 | 1800 watts | 1 | Y | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Advanced Chemistry Lab (1) |  | PSU Advanced Chemistry Lab | 430 | CLSB | 2 | 2N079 | EB121-2N079-1 | Balance, electronic, 121 grams x 1MG accuracy, top loading | Y | Mettler-Toledo | XP1203S | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Advanced Chemistry Lab (1) |  | PSU Advanced Chemistry Lab | 431 | CLSB | 2 | 2N079 | EB210-1N046-1 &2 | Balance, electronic, 210 grams x 0.01 gram accuracy | Y | Mettler-Toledo | XP204S | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Advanced Chemistry Lab (1) |  | PSU Advanced Chemistry Lab | 432 | CLSB | 2 | 2N079 | CB20-2N079-1-6 | Carboy, 40 L | Y | VWR | 89170-780 HDPE | [5 empty] | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Advanced Chemistry Lab (1) |  | PSU Advanced Chemistry Lab | 433 | CLSB | 2 | 2N079 | WPS-2N079-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 440 | CLSB | 2 | 2N047 |  | Shaker, stacked, incubated, no CO2 | Y | Thermo Scientific MaxQ 6000 | SHKE 6000-7 | Floor | 47 | 34 | 25 | Min | Open | Min | Open | 658 | 120 |  | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 443 | CLSB | 2 | 2N047 | LF-2N047-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | MF34SS-SARE-TS | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | Y | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 445 | CLSB | 2 | 2N047 | ULF-2N047-1 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 446 | CLSB | 2 | 2N047 | WB-2N047-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | [5 empty] | 120 | ?? | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 447 | CLSB | 2 | 2N047 | WPS-2N047-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 448 | CLSB | 2 | 2N047 |  | Incubator, 3.75 CF, no CO2 | Y | Thermo Fisher | Catalog #51028065, must have 2 chambers per incubator | Bench | 24 | 26 | 33 | Min | Open | Min | Open |  | 120 | 2.7 | 1 | N | Y | N | N | Y | std. | --- | ? | --- | --- | --- | --- | --- | N | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 449 | CLSB | 2 | 2N047 | EMS-2N047 | Monitoring & alarm system, lab equipment | Y | Thermo Fisher | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 450 | CLSB | 2 | 2N047 | UPP-2N047-1 & 2 | Pump, unispense peristaltic | Y | Wheaton | 13-683-25 | Bench | 12 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 451 | CLSB | 2 | 2N047 |  | pH meter, Bench top, pH/ISE multiparameter | Y | Thermo Scientific/ Orion | 13-645-542 STARA2146 | Bench | 12 | 12 | 6 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 452 | CLSB | 2 | 2N047 | EB121-2N047-1 & 2 | Balance, electronic, 121 grams x 1MG accuracy, top loading | Y | Mettler-Toledo | XP1203S | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Level 2 Biology Prep Lab (1) Dwg. LF102A |  | PSU General Biology Teaching Labs | 453 | CLSB | 2 | 2N047 | EB220-2N047-1 &2 | Balance, electronic, 220 grams x 0.001 gram accuracy | Y | Mettler-Toledo | ML203E | Bench | 7.76 | 15.68 | 3.84 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 459 | CLSB | 2 | 2N046, 2N036 |  | Microscope, compound, lab (25 per lab plus one for demonstration) | Y | Fisher Scientific Micromaster | 12-561-326 | Bench | [8 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 52 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 460 | CLSB | 2 | 2N046, 2N036 |  | Stereoscope, stereomaster zoom | Y | Fisher Scientific | 12-562-1, (10 per lab) | Bench | [8 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | [21 empty] | N | Y | OF/OI | NA | NA | 20 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 461 | CLSB | 2 | 2N046, 2N036 | DCI-2N046-1-12 DCI-2N036-1-12 | Labquest Data Collection Interface, Vernier | Y | Vernier | LABQ | Bench | M | M | M | --- | --- | --- | --- | --- | 120 | min | 1 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 462 | CLSB | 2 | 2N046, 2N036 | DCI-2N046-1-12 DCI-2N036-1-12 | Gas sensor, CO2 | Y | Vernier | CO2-BTA | Bench | M | M | M | --- | --- | --- | --- | --- | 120 | min | 1 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 463 | CLSB | 2 | 2N046, 2N036 | DCI-2N046-1-12 DCI-2N036-1-12 | Gas sensor, CO2 | Y | Vernier | O2-BTA | Bench | M | M | M | --- | --- | --- | --- | --- | 120 | min | 1 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 464 | CLSB | 2 | 2N046, 2N036 |  | Blender, laboratory | Y | Fisher Scientific | 14-509-10 | Bench | [8 empty] | 120 | min. | 1 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 465 | CLSB | 2 | 2N046, 2N036 |  | Indoor Grow Light Rack on wheels | Y |  |  | Bench | [8 empty] | 120 | ?? | 1 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 467 | CLSB | 2 | 2N046, 2N036 | CEN | Centrifuge | Y | Thermo Fisher Accuspin | 75002431 | Bench | 10 | 9 | 14 | --- | --- | --- | --- | --- | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 468 | CLSB | 2 | 2N046, 2N036 | WB5-2N046-1 thru 12 WB5-2N036-1 thru 12 | Water bath, 5 L | Y | Thermo Scientific | 15-460-S5Q, with SS cover | Bench | 14 | 11 | 8 | Open | Open | Open | Open |  | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 469 | CLSB | 2 | 2N046, 2N036 | WB-2N036 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | --- | --- | --- | --- | --- | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 470 | CLSB | 2 | 2N046, 2N036 | HP-1N046-1 thru 12 HP-1N036-1 thru 12 | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Isotemp | 11-200-49SH | Bench | 11 | 8 | 5 | Open | Open | Open | Open | --- | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 471 | CLSB | 2 | 2N046, 2N036 |  | Pipette, Pipettman Neo P200N | Y | Gilson | F144565 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 472 | CLSB | 2 | 2N046, 2N036 | P1000-2N046-1 thru 12 P1000-2N036-1 thru 12 | Pipette, Pipettman Neo P1000N | Y | Gilson | F144566 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 24 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 473 | CLSB | 2 | 2N046, 2N036 | EB410-2N046-1 thru 15 EB410-2N036-1 thru 15 | Balance, electronic, 410 grams x 0.01 gram accuracy, student grade | Y | Ohous Adventure Pro-Precision | 01-921-19 AV412-US | Bench | 9 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | PSU |  | Group III | Lab
PSU General Biology Teaching Labs (2) | 2N036 and 2N046 | PSU General Biology Teaching Labs | 474 | CLSB | 2 | 2N046, 2N036 | LS | Shaker, water bath, digital | Y | Thermo Scientific Max Q 7000 orbital | SHKE 7000-1CE, with universal platforms and clamps | Bench | 25 | 22 | 18 | M | Open | M | Open | 107 | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 480 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Microscope, compound | Y | Fisher Scientific | 12-561-326 | Bench | [8 empty] | 120 | min. | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 75 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 481 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Stereoscope, stereomaster zoom | Y | Fisher Scientific Micromaster | 12-562-1 | Bench | [8 empty] | 121 | min. | 2 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 482 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Labquest Interface, Vernier | Y | Vernier | LABQ | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 483 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Sensor, CO2 | Y | Vernier | CO2-BTA | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 484 | CLSB | 2 | 2N004, 2N016, 2N026 | EKG-2N004-1-12 EKG-2N016-1-12 EKG-2N026-1-12 | EKG Sensor | Y | Vernier | EKG-BTA | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 485 | CLSB | 2 | 2N004, 2N016, 2N026 | HD-2N004-1-12 HD-2N016-1-12 HD-2N026-1-12 | Dynamometer, hand | Y | Vernier | HD-BTA | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 486 | CLSB | 2 | 2N004, 2N016, 2N026 | CM-2N004-1-12 CM-2N016-1-12 CM-2N026-1-12 | Colorimeter | Y | Vernier | COL-BTA | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 487 | CLSB | 2 | 2N004, 2N016, 2N026 | UVSM-2N004-1 thru 3UVSM-2N016-1 thru UVSM-2N026-1 thru 33 | Spectronic 200 Spectrophotometer | Y | Thermo Scientific | S06240 | Bench |  |  |  | Open | Open | Open | Open |  | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 488 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Indoor Grow Light Rack on wheels | Y | ??? |  | Bench | [8 empty] | 120 | ??? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 490 | CLSB | 2 | 2N004, 2N016, 2N026 | CEN | Centrifuge | Y | Thermo Fisher Accuspin | 75002431 | Bench | 10 | 9 | 14 | --- | --- | --- | --- | --- | 120 | ?? | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 18 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 491 | CLSB | 2 | 2N004, 2N016, 2N026 | WB5-2N004-1 thru 12 WB5-2N016-1 thru 12 WB5-2N026-1 thru 12 | Water bath, 5 L | Y | Thermo Scientific | 15-460-S5Q, with SS cover | Bench | 14 | 11 | 8 | Open | Open | Open | Open | --- | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 492 | CLSB | 2 | 2N004, 2N016, 2N026 | WB-2N004-1 WB-2N016-1 WB-2N026-1 | Water bath, 20 L | Y | Thermo Fisher Isotemp 20 liter | Cat #2841 | Bench | 11 | 15 | 8 | --- | --- | --- | --- | --- | 120 | 600 watts | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 493 | CLSB | 2 | 2N004, 2N016, 2N026 | HP-2N004-1 thru 12 HP-2N016-1 thru 12 HP-2N026-1 thru 12 | Hot plates with magnetic stirrers, 7" x 7" | Y | Fisher Scientific Isotemp | 11-200-49SH | Bench | 11 | 8 | 5 | Open | Open | Open | Open | --- | 120 | 10 | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 494 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Pipette, Pipettman Neo P200N | Y | Gilson | F144565 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 495 | CLSB | 2 | 2N004, 2N016, 2N026 | P1000-2N004-1 thru 12 P1000-2N016-1 thru 12 P1000-2N026-1 thru 12 | Pipette, Pipettman Neo P1000N | Y | Gilson | F144566 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 36 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 496 | CLSB | 2 | 2N004, 2N016, 2N026 |  | Pipette, Pipettman Neo P10N | Y | Gilson | F144562 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 497 | CLSB | 2 | 2N004, 2N016, 2N026 | EB410-2N004-1 thru 15 EB410-2N016-1 thru 15 EB410-1N026-1 thru 15 | Balance, electronic, 410 grams x 0.01 gram accuracy, student grade | Y | Ohous Adventure Pro-Precision | 01-921-19 AV412-US | Bench | 9 | 12 | 12 | [5 empty] | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 18 | PSU |  | Group III | Lab
PSU Principals of Biology Teaching Labs (3) | Rooms 2N004, 2N016 and 2N026 | PSU Principals of Biology Teaching Labs | 498 | CLSB | 2 | 2N004, 2N016, 2N026 | LS | Shaker, water bath, digital | Y | Thermo Scientific Max Q 7000 orbital | SHKE 7000-1CE, with universal platforms and clamps | Bench | 25 | 22 | 18 | M | Open | M | Open | 107 | 120 | min | 1 | N | Y | N | N | Y | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | PSU |  | Group III | Lab
Inter Institutional Chemistry Labs (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 508 |  | 3N | 3N068 |  | Dewar, LN2 | Y | Technifab Products Inc. | TLN-050 (50 liter) | Floor | 18" dia. |  | 30" |  |  |  |  | 52 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
OCSSB Chemistry Lab (Dwg LF103B) |  | Dr. Gray - L3 West Labs & Lab Support | 529 | Dr. Gray | 3 | 3N074 | WPS-3N108-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Shared Glassware Wash (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 535 | CLSB | 3N | 3N108 | WPS-3N108-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
|  | CLSB - L3 Labs & Lab Support | 539 | CLSB | 3N | 3N072A |  | Biosafety cabinet, 6' | Y | The Baker Company | BioChemGARD BGC603A-HE, fully exhausted Class II A2 w/indirect venting by GC (replaces previous 6' chem fume hood) | Floor | 54 | 34 | 80 | Open | Open | Open | Open | 850 | 120 | 20 | 1 | Y | N | Y | N | Y | 5-20 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 12" | 664 | -1.5 | Yes | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Shared Solvent Room (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 543 | CLSB | 3N | 3N106 | LR-3N106-1 | Refrigerator Freezer 11651120 | Y | Thermo Scientific GP Series | 11651 120 3764A | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Shared Instrument Room (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 550 | CLSB | 3N | 3N098A | WPS-3N098A-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Shared Equipment Galley (Ultra Low Freezer Farm) (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 553 | CLSB | 3N | 3NH102 | ULF-3N067-1 thru 10 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 4 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | OSU |  | Group III | Lab
Shared Equipment Galley (Ultra Low Freezer Farm) (Dwg LF103B) |  | CLSB - L3 Labs & Lab Support | 554 | CLSB | 3N | 3NH102 | EMS-3N067 | Monitoring & alarm system, lab equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | OSU |  | Group III | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 556 | OSU | 3N | 3N095A | INC-3N095A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OSU |  | Group III | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 557 | OSU | 3N | 3N095A | BSC | Biosafety cabinet, 4' | Y | Baker SteriGARD | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 558 | OSU | 3N | 3N095A | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OSU |  | Group III | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 559 | OSU | 3N | 3N095A | WPS-3N095A-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 565 | OSU | 3N | 3N096A | INC-3N096A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OSU |  | Group III | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 566 | OSU | 3N | 3N096A | BSC | Biosafety cabinet, 4' | Y | Baker SteriGARD | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
OSU - Tissue Culture Labs (Dwg LF103A) |  | OSU - L3 Labs & Lab Support | 567 | OSU | 3N | 3N096A | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OSU |  | Group III | Lab
CLSB Tissue Culture Labs (Dwg LF103A) |  | CLSB Dr. Gray | 568 |  | 3N | 3N078A | INC-3N078A-1 &2 | Incubator, stacked, requires CO2, EXISTING | Y | ThermoForma | Haracell 1501 Trigas w roller base, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | BRB | New | 2 | OHSU |  | Group III | Lab
Shared Equipment Galley (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 573 | CLSB | 3N | 3NH078 | ULF-3NH078-1 thru 10 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco: refer to SV email 9/24/12 | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | OHSU |  | Group III | Lab
Shared Equipment Galley (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 574 | CLSB | 3N | 3NH078 | EMS-3NH078 | Monitoring & alarm system, lab equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 9 | OHSU |  | Group III | Lab
PSU - Tissue Culture Lab |  | CLSB - L3 Labs & Lab Support | 576 | CLSB | 3N | 3N093 | INC-3N093-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
PSU - Tissue Culture Lab |  | CLSB - L3 Labs & Lab Support | 577 | CLSB | 3N | 3N093 | BSC | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
PSU - Tissue Culture Lab |  | CLSB - L3 Labs & Lab Support | 578 | CLSB | 3N | 3N093 | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
Shared Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 579 | CLSB | 3N | 3NH094 | ULF-3NH094-1 thru 7 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 7 | OSU |  | Group III | Lab
Shared Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 580 | CLSB | 3N | 3NH094 | EMS-3NH094 | Monitoring & alarm system, lab equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OSU |  | Group III | Lab
Tissue Culture Labs (Dwg LF103A) |  | PSU - L3 Labs & Lab Support | 581 | PSU | 3N | 3N080A | INC-3N080A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
Tissue Culture Labs (Dwg LF103A) |  | PSU - L3 Labs & Lab Support | 582 | PSU | 3N | 3N080A | BSC | Biosafety cabinet, 4' | Y | Baker SteriGARD | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
Tissue Culture Labs (Dwg LF103A) |  | PSU - L3 Labs & Lab Support | 583 | PSU | 3N | 3N080A | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | PSU |  | Group III | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 585 | CLSB | 3N | 3N091A | INC-3N091A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 586 | CLSB | 3N | 3N091A | BSC | Biosafety cabinet, 4' | Y | Labconco | Class II Type A2 recirc | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 587 | CLSB | 3N | 3N091A | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 589 | CLSB | 3N | 3N090A | INC-3N092A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 590 | CLSB | 3N | 3N090A | BSC | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-1 | Group II |  | Lab
Inter-Institute Tissue Culture Labs (Dwg LF103A) |  | CLSB - L3 Labs & Lab Support | 591 | CLSB | 3N | 3N090A | LR | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Biomedical Labs (6 labs total on Level 3) Reference LF103A | Typical for rooms 3N080, 3N087 and 3N088 | Dr. Gray - L3 West Labs & Lab Support | 594 | CLSB | 3 | 3N080, 3N087, 3N088 | WPS-3N080-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-4 | Group II |  | Lab
Biomedical Labs (6 labs total on Level 3) Reference LF103A | OCSSB - Biomed lab | Dr. Gray - L3 West Labs & Lab Support | 600 | Dr. Gray | 3 | 3N082 | WPS-3N080-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-4 | Group II |  | Lab
Biomedical Labs (6 labs total on Level 3) Reference LF103A | Inter-Institute - Typical for rooms 3N090 and 3N091 | CLSB - L3 Labs & Lab Support | 605 | CLSB | 3N | 3N087A | WPS-3N080-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Shared Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 614 | CLSB | 3N | 3NH084 | EMS-3NH084 | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
OCSSB - Tissue Culture Lab (Dwg LF103A) |  | Dr. Gray - L3 West Labs & Lab Support | 616 | Dr. Gray | 3 | 3N082A | BSC | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | FFE-4 | Group II |  | Lab
OCSSB - Tissue Culture Labs (Dwg LF103A) |  | Dr. Gray - L3 West Labs & Lab Support | 626 |  | 3N | 3N087A | INC-3N087A-1 &2 | Incubator, stacked, requires CO2, plus stand | Y | ThermoForma | Heracell 150i, 13998 037 plus stand 12998 053 | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
OCSSB - Tissue Culture Labs (Dwg LF103A) |  | Dr. Gray - L3 West Labs & Lab Support | 627 | Dr. Gray | 3N | 3N087A | BSC | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | FFE-4 | Group II |  | Lab
OCSSB - Tissue Culture Labs (Dwg LF103A) |  | Dr. Gray - L3 West Labs & Lab Support | 628 | Dr. Gray | 3N | 3N087A | BSC | Biosafety cabinet, 6' | Y | Labconco Purifier Delta Series # 3460009 | Class II Type A2 recirc | Floor | 72 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. |  |  |  | X | [17 empty] | N | Y | OF/OI | NA | NA | 1 | FFE-4 | Group II |  | Lab
OCSSB - Tissue Culture Labs (Dwg LF103A) |  | Dr. Gray - L3 West Labs & Lab Support | 629 | Dr. Gray | 3N | 3N087A | LRF-3N082A-1 | Refrigerator/Freezer, manual defrost, residential grade | Y | Fisher Scientific Isotemp General Purpose: SV comment: Residential grade manual defrost 90/24/12 | 11 651 120 3764A, upright, 10 cu. Ft. | Floor | 31 | 31 | 60 | 3 | Open | 6 | Open | 300 | 120 | 5.8 | 1 | N | Y | N | N |  | Std. | --- | ? | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Shared Equipment Corridor - (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 649 | CLSB | 3N | 3NH086 | ULF-3NH086-1 thru 7 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco: refer to SV email 9/24/12 | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | OHSU |  | Group III | Lab
Shared Equipment Corridor - (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 651 | CLSB | 3N | 3NH102 |  | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco: refer to SV email 9/24/12 | UxF60086D | [44 empty] | 7 | OHSU |  | Group III | Lab
Shared Equipment Corridor - (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 652 | CLSB | 3N | 3NH102 |  | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 9 | OHSU |  | Group III | Lab
Shared Equipment Corridor - (Ultra Low Freezer Farm) |  | CLSB - L3 Labs & Lab Support | 653 | CLSB | 3N | 3NH086 | EMS-3NH086 | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 10 | OHSU |  | Group III | Lab
Inter-Institute Glassware & Media Prep (LF104B) |  | CLSB Level 4 Labs & Lab Support | 1260 |  | 4N | 4N110 | WPS-4N110-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1277 |  | 4N | 4NH107 | ULF-4NH094-1 thru 3 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1278 |  | 4N | 4NH107 | LF-4NH107-1 | Freezer, -20C, 34 CF upright, stainless steel with alarms | Y | Thermo Scientific GP series | MF34SS-SARE-TS | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | N | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1280 |  | 4N | 4NH107 | EMS | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Tissue Culture Labs (reference LF104A) | Typical for 3 labs: 4N082A, 4N096A and 4N094A | CLSB Level 4 Labs & Lab Support | 1296 | CLSB | 4N | 4N082A, 4N096A, 4N094A | INC-4N082A-1 INC-4N096A-1 INC-4N094-1 | Incubator, stacked, requires CO2, plus stand | Y | ThermoForma | Heracell 150i, Herai501 13998 037 3764A and stand 13998 053 | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 12 | OHSU |  | Group III | Lab
Tissue Culture Labs (reference LF104A) | Typical for 3 labs: 4N082A, 4N096A and 4N094A | CLSB Level 4 Labs & Lab Support | 1297 | CLSB | 4N | 4N082A, 4N096A, 4N094A | BSC-4N082A BSC-4N096A BSC-4N094 | Biosafety cabinet, 4' | Y | Labconco | Class II Type A2 recirc | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | FFE-1 | Group II |  | Lab
Tissue Culture Labs (reference LF104A) | Typical for 3 labs: 4N082A, 4N096A and 4N094A | CLSB Level 4 Labs & Lab Support | 1298 | CLSB | 4N | 4N082A, 4N096A, 4N094A | LR-4N082A LR-4N096A LR-4N094 | Refrigerator/Freezer, manual defrost, residential grade | Y | Thermo Scientific GP Series SV email 9/24/12: Reduce to residential grade | 11 651 120 3764A, upright, 10 cu. Ft. | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 3 | OHSU |  | Group III | Lab
Instrument Room |  | CLSB Level 4 Labs & Lab Support | 1302 | CLSB | 4N | 4N093 | WPS-4N093-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-1 | Group II |  | Lab
Inter-Institute Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1307 | CLSB | 4N | 4NH094 | ULF-4NH094-1 thru 10 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 10 | OHSU |  | Group III | Lab
Inter-Institute Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1308 | CLSB | 4N | 4NH094 | LN2-4NH094-1 LN2-4NH094-2 | Dewar, LN2 | Y | Technifab Products Inc. | TLN-050 (50 liter) | Floor | 18" dia. |  | 30" |  |  |  |  | 52 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |  |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
Inter-Institute Equipment Corridor (Ultra Low Freezer Farm) |  | CLSB Level 4 Labs & Lab Support | 1310 | CLSB | 4N | 4NH094 | EMS-4NH094 | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Tissue Culture Labs | Typical for 4 TC labs: 4N097A, 4N081A, 4N081B and 4NH095 | CLSB Level 4 Labs & Lab Support | 1319 | CLSB | 4N | 4N097A, 4N081A, 4N081B,4N095 | INC-4N097A-1 INC-4N081A-1 INC-4N081B-1 INC-4NH094-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | Heracel 13 998 051, sold per pair | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 16 16 | 1 1 | Y Y | N N | Y Y | N N | N N | Std. | --- | Y Y | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 16 | OHSU |  | Group III | Lab
Tissue Culture Labs | Typical for 4 TC labs: 4N097A, 4N081A, 4N081B and 4NH095 | CLSB Level 4 Labs & Lab Support | 1320 | CLSB | 4N | 4N097A, 4N081A, 4N081B,4N095 | BSC-4N097A BSC-4N081A BSC-4N081B BSC-??? | Biosafety cabinet, 4' | Y | Labconco | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 8 | FFE-1 | Group II |  | Lab
Tissue Culture Labs | Typical for 4 TC labs: 4N097A, 4N081A, 4N081B and 4NH095 | CLSB Level 4 Labs & Lab Support | 1321 | CLSB | 4N | 4N097A, 4N081A, 4N081B,4N095 | LR-4N097A LR-4N081A LR-4N081B LR-??? | Refrigerator/Freezer, manual defrost | Y | Thermo Scientific GP Series | 11 651 120 3764A, lab grade, 20 CF | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 4 | OHSU |  | Group III | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | Skourtes Research | 1331 |  | 4N | 4N074 | ULF-4N074-1 thru 14 | Freezer, -80C, Ultra Low 28.8 Cu. Ft. | Y | Thermo Scientific Revco | UxF60086D | Floor | 44 | 38 | 78 | M | Open | M | Open |  | 208 | 9 | 1 | Y | N | Y | N | N | 6-20P |  | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 14 | OHSU |  | Group III | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | Skourtes Research | 1332 |  | 4N | 4N074 |  | Dewar, LN2 | Y | Technifab Products Inc. | TLN-050 (50 liter) | Floor | 18" dia. |  | 30" |  |  |  |  | 52 | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
Equipment Corridor (Ultra Low Freezer Farm) |  | Skourtes Research | 1334 |  | 4N | 4N074 | EMS | Monitoring & alarm system, critical equipment | N | Thermo Scientific | Smart-Vue Wireless | Special | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Tissue Culture Labs (reference LF104A) | Typical for 2 labs: 4N077A and 4N078A | Skourtes Research | 1343 |  | 4N |  | BSC-4N077A BSC-4N078A | Biosafety cabinet, 4' | Y | Baker SteriGARD | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | FFE-1 | Group II |  | Lab
| Soiled Holding & Manual Cleaning Area | SOD Central Sterile | 1354 | SKT | 5 | 5N095 | UC | Cleaner, ultrasonic | N | TBD | TBD | Bench | [8 empty] | 120 | min | 1 | [9 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
Tissue Culture Labs | Typical TC lab equipment for rooms 6N102A, 6N096A, 6N088A, 6N088B, 6N095A, 6N087A and 6N084A. | SKT Level 6 Labs & Lab Support | 1422 | SKT-R | 6 | 6N102A, 6N096A, 6N088A, 6N088B,6N095A, 6N087A, 6N084A | INC-6N102A-1 INC-6N096A-1 INC-6N088A-1 INC-6N088B-1 INC-6N095A-1 INC-6N087A-1 INC-6N084A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 18 18 | 1 1 | Y Y | N | X X | N | N |  | --- |  | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 16 | OHSU |  | Group III | Lab
Tissue Culture Labs | Typical TC lab equipment for rooms 6N102A, 6N096A, 6N088A, 6N088B, 6N095A, 6N087A and 6N084A. | SKT Level 6 Labs & Lab Support | 1423 | SKT-R | 6 | 6N102A, 6N096A, 6N088A, 6N088B,6N095A, 6N087A, 6N084A | BSC- | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 14 | FFE-2 | Group II |  | Lab
Tissue Culture Labs | Typical TC lab equipment for rooms 6N102A, 6N096A, 6N088A, 6N088B, 6N095A, 6N087A and 6N084A. | SKT Level 6 Labs & Lab Support | 1424 | SKT-R | 6 | 6N102A, 6N096A, 6N088A, 6N088B,6N095A, 6N087A, 6N084A | LR-6N102A-1 LR-6N096A-1 LR-6N088A-1 LR-6N088B-1 LR-6N095A-1 LR-6N087A-1 LR-6N084A-1 | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 7 | OHSU |  | Group III | Lab
Tissue Culture Labs | Tissue Culture | SKT Level 6 Labs & Lab Support | 1429 | SKT-R | 6 | 6N099A | INC-6N099A-1 | Incubator, stacked, requires CO2 | Y | ThermoForma | SteriCult 3310, 2 units | Floor | 43 | 27 | 40 80 | M | Open |  | Open |  | 120 120 | 18 18 | 1 1 | Y Y | N | X X | N | N |  | --- |  | --- | --- | --- | --- | --- | X X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Tissue Culture Labs | Tissue Culture | SKT Level 6 Labs & Lab Support | 1430 | SKT-R | 6 | 6N099A | BSC- | Biosafety cabinet, 4' | Y | Labconco Purifier Delta Series # 3440009 | Class II Type A2 recirc, fully exhausted Class II A2 w/indirect venting by GC | Floor | 48 | 31 | 99 | M | Open | M | Open |  | 120 | 18 | 1 | Y | N | Y | N | N | Std. | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | FFE-2 | Group II |  | Lab
Tissue Culture Labs | Tissue Culture | SKT Level 6 Labs & Lab Support | 1431 | SKT-R | 6 | 6N099A | LR-6N099A-1 | Refrigerator, w/alarm & chart recorder | Y | Thermo Scientific GP Series | MR30SS-SARE-TS, lab grade, 27 CF, solid SS door, SS interior | Floor | 30 | 32 | 78 | M | open | M | open | 380 | 120 | 12? | 1 | Y | N | Y | N | Y | Std. | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Bio-Chemistry Research Labs (2 bio-chem labs on Level 6) (LF106B) | Typical for rooms 6N096, 6N098, 6N100 and 6N102 | SKT Level 6 Labs & Lab Support | 1437 | SKT-R | 6 | 6N096, 6N098, 6N100, 6N102 | WPS | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-2 | Group II |  | Lab
Biomedical Research Labs (8 labs total on Level 6) | Typical for biomed labs 6N095, 6N099, 6N088, 6N087, 6N086, 6N084, 6N083, and 6N085 | SKT Level 6 Labs & Lab Support | 1447 | SKT-R | 6 | 6N095,6N099,6N088,6N087,6N086,6N084,6N083,6N085 | WPS-6N095-1 WPS-6N099-1 WPS-6N088-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 3 | FFE-2 | Group II |  | Lab
Instrument Room (LF106A) |  | SKT Level 6 Labs & Lab Support | 1454 | SKT-R | 6 | 6N094 | WPS-6N094-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-2 | Group II |  | Lab
Biomedical Lab Support / Instrument Room |  | SKT Level 6 Labs & Lab Support | 1471 | SKT-R | 6 | 6N091 | WPS-6N091-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-2 | Group II |  | Lab
OHSU Glassware Wash and Dirty Media Prep (LF106A) |  | SKT Level 6 Labs & Lab Support | 1487 | SKT-R | 6 | 6N081 | WPS-6N081-1 | Water purification system (polisher) | Y | Barnstead GenPure xCAD UV/UF-TOC w/ bench mtd. dispenser | PN 50131298 | Wall | 15 | 13 | 24 | Min. | Open | Min. | open |  | 120 | 100 VA | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | N | Y | OF/VI | NA | NA | 1 | FFE-2 | Group II |  | Lab
Animal Holding |  | SKT Level 6 Labs & Lab Support | 1495 | SKT-R | 6 | 6N080A |  | Chair, sanitizable | N | TBD | TBD | [39 empty] | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
OHSU Procedure Room |  | SKT Level 6 Labs & Lab Support | 1498 | SKT-R | 6 | 6N080AA | BSC- | Biosafety cabinet / Animal Transfer Station, on casters | Y | Labconco Purifier Delta Series # 3440009 | SG403A HE ATS | Floor | 64 | 31 | 80 | Min. | Open | Min. | Open | 946 | 120 | 20 | 1 | Y | N | Y | N | Y | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | FFE-2 | Group II |  | Lab
IGP - Model Trimming / Clinical Lab |  | SOD Level 8 IGP | 1558 | SKT | 8 | 8N084 | BB-1 | Bunsen burner with shield | N | Almor | TBD | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 0 | --- | --- | --- | --- | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | NA | NA | 8 | OHSU |  | Group III | Lab
IGP - Model Trimming / Clinical Lab |  | SOD Level 8 IGP | 1570 | SKT | 8 | 8N084 | MIC-1 | Microscope, w/ individual lamp | Y | Harris (Meiji optics) | EMT-L | Bench | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 120 | min | 1 | N | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
IGP - Clean Dispensary & Soiled Holding |  | SOD Level 8 IGP | 1575 | SKT | 8 | 8N099 |  | Monitor arm | N | Ergotron?? | TBD, wall mounted | Wall |  |  |  | open | open | open | open |  | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
IGP - Model Trimming / Clinical Lab |  | SOD Level 9 IGP | 1640 | SKT | 9 |  | BB-1 | Bunsen burner | N | TBD | TBD | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 0 | --- | --- | --- | --- | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | NA | NA | 7 | OHSU |  | Group III | Lab
OHSU - SOD Dental Simulation | Dental Wet Lab | SOD Dental Simulation | 1716 | SKT | 10 | 10N010 | BO-1 | Burnout Ovens | Y | Vulcan | A550 | Bench | 16 | 16 | 16 | min | open | min | open | 0 | 120 | 13 | 1 | Y | Y | N | N | N | Std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 6 | OHSU |  | Group III | Lab
OHSU - SOD Dental Simulation | Dental Wet Lab | SOD Dental Simulation | 1717 | SKT | 10 | 10N010 | 0 | Torch, natural gas | N | 0 | 0 | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 5 | OHSU |  | Group III | Lab
OHSU - SOD Dental Simulation | Dental Dry Lab | SOD Dental Simulation | 1723 | SKT | 10 | 10N014 |  | Microscope, w/ individual lamp | Y | Harris (Meiji optics) | EMT-L | Bench | [8 empty] | 120 | min | 1 | Y | Y | N | N | N | Std. | [21 empty] | N | Y | OF/OI | NA | NA | 12 | OHSU |  | Group III | Lab
Pediatrics | Peds Graduate Support Lab | SOD Pediatrics | 1829 | SKT | 11 | 11N100 | BB-1 | Bunsen burner | N | TBD | TBD | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 0 | --- | --- | --- | --- | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Oral Surgery & Sedation | Surgery & Sedation Procedure Rooms (Typical for 3 rooms) | SOD Oral Surgery & Sedation | 1835 | SKT | 11 | 11N086, 11N090, 11N091 |  | Reservoir, vacuum | Y | TBD | TBD | Floor | [38 empty] | 0 | 0 | 0 | 0 | 0 | 3 | FFE-2 | Group II |  | Lab
Oral Surgery & Sedation | Oral Surgery - Pre-Doc Operatories (Typical of 6 rooms) | SOD Oral Surgery & Sedation | 1863 | SKT | 11 | 11N072, 11N073, 11N074, 11N075, 11N076, 11N077 |  | Reservoir, vacuum w/hoses | Y | TBD | TBD | Floor | [38 empty] | N | Y | OF/OI | NA | NA | 6 | FFE-2 | Group II |  | Lab
Oral Surgery & Sedation | Clean Equipment Storage, Pharm Dispensary, Soiled Holding & SAT Instrument Processing | SOD Oral Surgery & Sedation | 1904 | SKT | 11 | 11N082, 11N084B |  | Refrigerator, under counter, scientific, manual defrost | N | Thermo Fisher | TBD | Floor | 24 | 24 | 34.5 | min. | open | 4' | min. |  | 120 | std. | 1 | N | Y | N | N | N | std. | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | N | Y | OF/OI | NA | NA | 1 | OHSU |  | Group III | Lab
Periodontics | Perio/Endo Support Lab | SOD Periodontics | 1967 | SKT | 11 | 11N057 | BB-1 | Bunsen burner | N | TBD | TBD | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 0 | --- | --- | --- | --- | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
AEGD & CE | AEGD/CE - Model Trimming/Clinical Support Lab | SOD AEGD/CE | 2073 | SKT | 12 | 12N122 | BB-1 | Bunsen burner ( Electric Waxers per 4/4 list) | N | TBD | TBD | Bench | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | 0 | --- | --- | --- | --- | --- | --- | --- | X | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | Y | N | OF/OI | NA | NA | 2 | OHSU |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2161 | CLSB | TBD | TBD |  | Beaker |  | VWR | 89000-202 |  | 150 mL | [42 empty] | 1512 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2162 | CLSB | TBD | TBD |  | Beaker |  | VWR | 89000-204 |  | 250 mL | [42 empty] | 1512 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2163 | CLSB | TBD | TBD |  | Beaker |  | VWR | 89000-206 |  | 400 mL | [42 empty] | 1512 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2164 | CLSB | TBD | TBD |  | Beaker |  | VWR | 89000-210 |  | 800 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2165 | CLSB | TBD | TBD |  | Erlenmeyer Flask |  | VWR | 29136-048 |  | 125 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2166 | CLSB | TBD | TBD |  | Erlenmeyer Flask |  | VWR | 29136-060 |  | 250 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2167 | CLSB | TBD | TBD |  | Volumetric Flask |  | VWR | 89004-066 |  | 50 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2168 | CLSB | TBD | TBD |  | Volumetric Flask |  | VWR | 89004-068 |  | 100 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2169 | CLSB | TBD | TBD |  | Graduated Cylinder |  | VWR | 89001-142 |  | 10 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2170 | CLSB | TBD | TBD |  | Graduated Cylinder |  | VWR | 89001-144 |  | 25 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2172 | CLSB | TBD | TBD |  | Test Tube Rack |  | Fisher | S95921 |  | polypropylene, blue holds 6 tubes | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2173 | CLSB | TBD | TBD |  | Test Tube |  | VWR | 89000-842 |  | 13x100mm | [42 empty] | 4536 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2174 | CLSB | TBD | TBD |  | Test Tube |  | VWR | 89000-848 |  | 18x150mm | [42 empty] | 3024 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2175 | CLSB | TBD | TBD |  | Test Tube |  | VWR | 89001-432 |  | 25x200mm | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2176 | CLSB | TBD | TBD |  | Seriological Pipet |  | VWR | 93000-680 |  | 10mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2177 | CLSB | TBD | TBD |  | Wash Bottle |  | VWR | 414004-226 |  | 250 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2178 | CLSB | TBD | TBD |  | Crucible Tongs |  | Fisher | S50256 |  | 9 inch | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2179 | CLSB | TBD | TBD |  | Polypropylene Short Stem Funnel |  | VWR | 470145-358 |  | Top ID, mm: 75, Stem Length, mm: 75 | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2180 | CLSB | TBD | TBD |  | Stir Bars |  | Fisher | 14-512-126 |  | 9.5dia x 25.4mm | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2181 | CLSB | TBD | TBD |  | Scoopula, Lab Scoop |  | Fisher | S02602 |  | 6" | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2182 | CLSB | TBD | TBD |  | test tube holder/clamp |  | Fisher | S49062A |  | 5" | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2183 | CLSB | TBD | TBD |  | Medicine Dropper |  | Fisher | S18765 |  | 2mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2184 | CLSB | TBD | TBD |  | Medicine Dropper Bulb |  | VWR | 15001-362 |  | 1 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2185 | CLSB | TBD | TBD |  | Pasteur Pipets |  | Fisher | 63A53 |  | 5/75" | [42 empty] | 3024 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2186 | CLSB | TBD | TBD |  | Pipet Bulbs |  | Fisher | S32324 |  | 1 mL | [42 empty] | 756 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2187 | CLSB | TBD | TBD |  | Weigh Boats |  | Fisher | 13-735-742 |  | small | [42 empty] | 3024 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2188 | CLSB | TBD | TBD |  | Support Stand w/ Rod & Ring |  | VWR | 470005-866 |  | 13 cm x 20 cm | [42 empty] | 96 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2189 | CLSB | TBD | TBD |  | Clamp Holder, Aluminum |  | VWR | 21572-556 |  | N/A | [42 empty] | 120 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2190 | CLSB | TBD | TBD |  | Double-Buret Clamps |  | Fisher | S90318 |  | N/A | [42 empty] | 96 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2191 | CLSB | TBD | TBD |  | Three-Prong Extension Clamps |  | Fisher | 05-769-6Q |  | Length: 6.5 in. Extension arm: 4.25 in | [42 empty] | 144 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2192 | CLSB | TBD | TBD |  | Three-Prong Extension Clamps |  | Fisher | 05-769-7Q |  | Length: 8.75 in. Extension arm: 5.12 in. | [42 empty] | 72 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2193 | CLSB | TBD | TBD |  | Basic Stirring/Heat Plates |  | Fisher | 11-100-16SH |  | 4"x 4" | [42 empty] | 72 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2194 | CLSB | TBD | TBD |  | Well Plates |  | Fisher | 07-200-82 |  | 12 wells | [42 empty] | 72 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2195 | CLSB | TBD | TBD |  | Filter Flasks (stopper number 7) |  | Fisher | 10-180E |  | 500mL, 10mm tubulation OD | [42 empty] | 72 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2196 | CLSB | TBD | TBD |  | Graduated Cylinder |  | VWR | 89001-148 |  | 100 mL | [42 empty] | 144 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2197 | CLSB | TBD | TBD |  | Pipet Filler |  | VWR | 53497-055 |  | VWR | [42 empty] | 72 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2198 | CLSB | TBD | TBD |  | Vernier LabPRO |  | Vernier | LABPRO |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2199 | CLSB | TBD | TBD |  | Temperature Probe |  | Vernier | TMP-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2200 | CLSB | TBD | TBD |  | Colorimeter |  | Vernier | COL-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2201 | CLSB | TBD | TBD |  | Conductivity Probe |  | Vernier | CON-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2202 | CLSB | TBD | TBD |  | pH Sensor |  | Vernier | PH-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2203 | CLSB | TBD | TBD |  | Gas Pressure Sensor |  | Vernier | GPS-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2204 | CLSB | TBD | TBD |  | Voltage Probe |  | Vernier | VP-BTA |  | N/A | [42 empty] | 24 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2205 | CLSB | TBD | TBD |  | Spectrometers |  | Vernier | SPRT-UV-VIS |  | N/A | [42 empty] | 12 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2206 | CLSB | TBD | TBD |  | Spectrophotometer BIOwave II |  | BIOEXPRESS | 80-3003-75 |  | Biochrom BIOWAVE II | [42 empty] | 6 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2207 | CLSB | TBD | TBD |  | Macrokjeldahl Heaters |  | Cole Parmer | MQ3866B/EX1 |  | 6 well Bibby/Electrotherm | [42 empty] | 4 |  |  | Group III | Lab
|  | PSU General Chemistry Teaching Labs | 2208 | CLSB | TBD | TBD |  | University Meter Kit |  | Accumet | 13-636-AB15E | [44 empty] | 4 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2209 | CLSB | TBD | TBD |  | Simple squamous epithelium sec. (Bowman's capsule) |  | Ward's | 93 W 3004 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2210 | CLSB | TBD | TBD |  | Frog Skin, c.s. |  | Triarch Inc. | ZL 7-24 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2211 | CLSB | TBD | TBD |  | Skin sole of foot (human), section |  | Ward's | HI 1-2 | [44 empty] | 30 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2212 | CLSB | TBD | TBD |  | Simple cuboidal epithelium thin sec. H&E |  | Ward's | 96 W 3024 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2213 | CLSB | TBD | TBD |  | Salivary Glands Serous, Mucous, and Mixed; sections |  | Ward's | HK 4-51 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2214 | CLSB | TBD | TBD |  | Epithelium columnar; goblet cells, sections |  | Ward's | HA 3-33 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2215 | CLSB | TBD | TBD |  | Submaxillary gland primate sec. |  | Ward's | 93 W 8552 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2216 | CLSB | TBD | TBD |  | Pseudostratified ciliated columnar epithelium |  | Triarch Inc. | 2--72 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2217 | CLSB | TBD | TBD |  | Mammal Transitional Epi sec. |  | Carolina | 31-2462 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2218 | CLSB | TBD | TBD |  | Areolar tissue |  | Triarch Inc. | HB 3-1 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2219 | CLSB | TBD | TBD |  | White Fibrous Tissue, tendon: l.s. (Human) |  | Triarch Inc. | 15-75 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2220 | CLSB | TBD | TBD |  | Reticular tissue sec. m&nfr |  | Ward's | 93 W 3236 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2221 | CLSB | TBD | TBD |  | Adipose Tissue Fat cells (Human), section |  | Triarch Inc. | HB 4-1 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2222 | CLSB | TBD | TBD |  | Mammal Hyaline Cartilage sec. |  | Carolina | 31-2886 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2223 | CLSB | TBD | TBD |  | Mammal Elastic Cartilage |  | Carolina | 31-2910 (H690) | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2224 | CLSB | TBD | TBD |  | Cartilage Fibrous, section |  | Triarch Inc. | HB 7-14 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2225 | CLSB | TBD | TBD |  | Blood human (Giemsa Stain), smear |  | Triarch Inc. | HC 1-4 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2226 | CLSB | TBD | TBD |  | Mammal Compact Bone Ground, c.s. |  | Carolina | 31-2964 (G6325) | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2227 | CLSB | TBD | TBD |  | Muscle types smooth, cardiac & skeletal sec. H&E |  | Ward's | 93 W 3513 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2228 | CLSB | TBD | TBD |  | Skeletal muscle human c.s & l.s. |  | Ward's | 93 W 6270 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2229 | CLSB | TBD | TBD |  | Cardiac muscle human |  | Carolina | H 7075 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2230 | CLSB | TBD | TBD |  | Smooth muscle intestine c.s. & l.s |  | Ward's | 93 W 3520 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2231 | CLSB | TBD | TBD |  | Mammal Neuron Motor Nerve Cells smear |  | Carolina | 31-3570 (H1660) | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2232 | CLSB | TBD | TBD |  | Medullated nerve c.s. & l.s. |  | Ward's | 93 W 3651 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2233 | CLSB | TBD | TBD |  | Cerebellum Mammal silver |  | Carolina | H 1510 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2234 | CLSB | TBD | TBD |  | Spinal cord mammal, 4 levels cervical, thoracic, lumbar, sacral W&NFR |  | Ward's | 93 W 3707 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2235 | CLSB | TBD | TBD |  | Mammal Spinal Ganglion and Nerve l.s. |  | Carolina | 31-3792 (H1592) | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2236 | CLSB | TBD | TBD |  | Eye monkey l.s. |  | Ward's | 93 W 8290 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2237 | CLSB | TBD | TBD |  | Cochlea guinea pig l.s. H&E |  | Ward's | 93 W 3775 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2238 | CLSB | TBD | TBD |  | Thyroid parathyroid sec. |  | Carolina | H 4300 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2239 | CLSB | TBD | TBD |  | Pancreas exocytotic vessicles, mammal sec. H&E |  | Ward's | 93 W 4600 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2240 | CLSB | TBD | TBD |  | Thymus Hassall's corpuscles c.s. H&E |  | Ward's | 93 W 4122 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2241 | CLSB | TBD | TBD |  | Pineal body human sec. |  | Ward's | 93 W 6616 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2242 | CLSB | TBD | TBD |  | Mammal hypophysis pituitary gland sag. Sec. H&E |  | Ward's | 93 W 4334 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2243 | CLSB | TBD | TBD |  | Adrenal gland human sec. |  | Ward's | 93 W 6610 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2244 | CLSB | TBD | TBD |  | Trachea ciliated epithelium c.s. H&E |  | Ward's | 93 W 4871 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2245 | CLSB | TBD | TBD |  | Olfactory epithelim mammal sec. H&E |  | Ward's | 93 W 3787 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2246 | CLSB | TBD | TBD |  | Lung human sec. |  | Ward's | 93 W 6920 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2247 | CLSB | TBD | TBD |  | Human Pulmonary Emphysema, sec. |  | Carolina | 31-7804 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2248 | CLSB | TBD | TBD |  | Tooth in situ l.s. |  | Ward's | 93 W 4454 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2249 | CLSB | TBD | TBD |  | Tongue foliate papilla (for taste buds), section |  | Ward's | HK 2-26 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2250 | CLSB | TBD | TBD |  | Parotid gland human sec. |  | Ward's | 93 W 6730 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2251 | CLSB | TBD | TBD |  | Salivary Glands Submandibular (submaxillary), sec. |  | Triarch Inc. | HK 4-2 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2252 | CLSB | TBD | TBD |  | Sublingual gland human sec. |  | Ward's | 93 W 6734 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2253 | CLSB | TBD | TBD |  | Mammal Salivary Glands Composite, sec. |  | Carolina | 31-4986 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2254 | CLSB | TBD | TBD |  | Human liver section |  | Carolina | 315388 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2255 | CLSB | TBD | TBD |  | Human Gallbladder, Body sec. |  | Carolina | 31-5424 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2256 | CLSB | TBD | TBD |  | Esophagus c.s. H&E |  | Ward's | 93 W 4498 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2257 | CLSB | TBD | TBD |  | Cardiac stomach mammal sec. H&E |  | Ward's | 93 W 4510 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2258 | CLSB | TBD | TBD |  | Stomach mammal sec. H&E |  | Ward's | 93 W 4508 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2259 | CLSB | TBD | TBD |  | Pyloric stomach primate sec. |  | Ward's | 93 W 8489 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2260 | CLSB | TBD | TBD |  | Duodenum mammal c.s. H&E |  | Ward's | 93 W 4526 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2261 | CLSB | TBD | TBD |  | Jejunum mammal c.s. H&E |  | Ward's | 93 W 4530 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2262 | CLSB | TBD | TBD |  | Ileum Peyer's patches, microvilli, desmosomes c.s. H&E |  | Ward's | 93 W 4534 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2263 | CLSB | TBD | TBD |  | Appendix human: c.s. |  | Triarch Inc. | HK 8-3 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2264 | CLSB | TBD | TBD |  | Colon human sec. |  | Ward's | 93 W 6790 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2265 | CLSB | TBD | TBD |  | Rectum Anal opening c.s. HK 8-25 |  | Triarch Inc. | TBD | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2266 | CLSB | TBD | TBD |  | Kidney (Human) section |  | Triarch Inc. | HL 1-21 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2267 | CLSB | TBD | TBD |  | Mammal kidney sec. H&E |  | Ward's | 93 W 5234 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2268 | CLSB | TBD | TBD |  | Mammal urinary bladder collapsed c.s. H&E |  | Ward's | 93 W 5254 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2269 | CLSB | TBD | TBD |  | Mammal ureter c.s. H&E |  | Ward's | 93 W 5242 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2270 | CLSB | TBD | TBD |  | Urethra mammal female c.s. h&e |  | Ward's | 93 W 5246 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2271 | CLSB | TBD | TBD |  | Rat ovary oogenesis |  | Ward's | [45 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2272 | CLSB | TBD | TBD |  | Mammal ovary corpus luteum of ovulation sec. mal |  | Ward's | 93 W 5536 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2273 | CLSB | TBD | TBD |  | Ovary corpus albicans human sec. |  | Ward's | 93 W 7272 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2274 | CLSB | TBD | TBD |  | Uterus human progravid phase sec. |  | Ward's | 93 W 7306 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2275 | CLSB | TBD | TBD |  | Uterus human foll. Phase sec. |  | Ward's | 93 W 7303 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2276 | CLSB | TBD | TBD |  | Uterus human menstrual phase sec. |  | Ward's | 93 W 7314 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2277 | CLSB | TBD | TBD |  | Vagina (monkey) sec. |  | Ward's | 93 W 8772 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2278 | CLSB | TBD | TBD |  | Oviduct, fimbria: sec. (human) |  | Triarch Inc. | HN 2-22 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2279 | CLSB | TBD | TBD |  | Testis Spermatogenesis(iron hematoxylin), section |  | Triarch Inc. | HM 1-2 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2280 | CLSB | TBD | TBD |  | Epididymis priamte sec. |  | Ward's | 93 W 8723 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2281 | CLSB | TBD | TBD |  | Human Penis sec. |  | Ward's | 31-6566 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2282 | CLSB | TBD | TBD |  | Seminal vesicle human sec. |  | Ward's | 93 W 7228 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2283 | CLSB | TBD | TBD |  | Prostate, young adult sec. (Human) |  | Triarch Inc. | HM 5-2 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2284 | CLSB | TBD | TBD |  | Vas deferens: c.s. (Human) |  | Triarch Inc. | HM 4-11 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2285 | CLSB | TBD | TBD |  | Plastic slide boxes |  | Carolina | RE-634217 | [44 empty] | 12 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2286 | CLSB | TBD | TBD |  | brain |  | Carolina | RE-566781 | [44 empty] | 7 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2287 | CLSB | TBD | TBD |  | brain ventricles |  | Denoyer-Geppert | A410 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2288 | CLSB | TBD | TBD |  | sheep brain biosmount |  | Carolina | RE-265310 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2289 | CLSB | TBD | TBD |  | Unisex torso |  | Carolina | AWB19 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2290 | CLSB | TBD | TBD |  | lung |  | Carolina | RE-566917 | [44 empty] | 10 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2291 | CLSB | TBD | TBD |  | pancreas/gallbladder/spleen |  | Anatomy warehouse | AW333 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2292 | CLSB | TBD | TBD |  | liver with gallbladder |  | Anatomy warehouse | AW330 | [44 empty] | 4 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2293 | CLSB | TBD | TBD |  | kidney |  | Anatomy warehouse | AW325 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2294 | CLSB | TBD | TBD |  | Giant inner middle and outer ear |  | Anatomy warehouse | AWE10 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2295 | CLSB | TBD | TBD |  | inner ear |  | Carolina | RE-566960 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2296 | CLSB | TBD | TBD |  | ear ossicles |  | Denoyer-Geppert | DGA32 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2297 | CLSB | TBD | TBD |  | male half pelvis |  | Anatomy warehouse | AWMAH11 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2298 | CLSB | TBD | TBD |  | female half pelvis |  | Anatomy warehouse | AWMAH10 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2299 | CLSB | TBD | TBD |  | Arm musculature model |  | Sargent-welch | WL1770-09 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2300 | CLSB | TBD | TBD |  | leg muscles |  | Sargent-welch | WL1770-10 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2301 | CLSB | TBD | TBD |  | flexible vertebral column |  | Carolina | RE-247573 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2302 | CLSB | TBD | TBD |  | Spinal cord |  | Denoyer-Geppert | A65 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2303 | CLSB | TBD | TBD |  | Bone model |  | Ward's | RE-810660 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2304 | CLSB | TBD | TBD |  | Eye in orbit |  | Sargent-welch | WLH85245 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2305 | CLSB | TBD | TBD |  | Human eye model |  | Carolina | RE-566945A | [44 empty] | 3 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2306 | CLSB | TBD | TBD |  | Heart model |  | Carolina | RE-566810 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2307 | CLSB | TBD | TBD |  | Disarticulated skeleton |  | Skulls unlimited | WSP-14 | [44 empty] | 8 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2308 | CLSB | TBD | TBD |  | articulated arm |  | Carolina | RE-247710 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2309 | CLSB | TBD | TBD |  | leg w/ .5 pelvis |  | Carolina | RE-247930 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2310 | CLSB | TBD | TBD |  | Skulls calvaraium |  | Skulls unlimited | WS-BC-212 | [44 empty] | 6 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2311 | CLSB | TBD | TBD |  | Magnetic Adult Human skull |  | Shopanatomical.com | EZ-Z4701 | [44 empty] | 3 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2312 | CLSB | TBD | TBD |  | skull Didactic version (colored bones) |  | Anatomy warehouse | AWA291 | [44 empty] | 4 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2313 | CLSB | TBD | TBD |  | Functional larynx |  | Anatomy warehouse | AWG20 | [44 empty] | 2 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2314 | CLSB | TBD | TBD |  | 4-part half skull |  | Anatomy warehouse | AWA280 | [44 empty] | 4 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2315 | CLSB | TBD | TBD |  | Dissecting pans |  | Carolina | 629004 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2316 | CLSB | TBD | TBD |  | Scalpels |  | Carolina | 626520 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2317 | CLSB | TBD | TBD |  | Scalpel blades |  | Carolina | 626631 | [44 empty] | 10 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2318 | CLSB | TBD | TBD |  | Scissors |  | Carolina | 621910 | [44 empty] | 20 |  |  | Group III | Lab
|  | PSU Anatomy & Physiology Labs | 2319 | CLSB | TBD | TBD |  | Dissecting pins |  | Bio-Corporation | TP0120R | [44 empty] | 3 |  |  | Group III | Lab
|  | PSU Micro/Cell/Recomb DNA Labs | 2320 | CLSB |  | 1N044, IN046, IN076 |  | Shaking water bath, 20L |  | Thermo Scientific | S05981 | [44 empty] | 3 |  |  | Group III | Lab
|  | PSU Micro/Cell/Recomb DNA Labs | 2321 | CLSB |  | 1N044, IN046, IN077 |  | Pipette, Pipettman Neo P200N |  | Gilson | F144565 | [44 empty] | 36 |  |  | Group III | Lab
|  | PSU Micro/Cell/Recomb DNA Labs | 2322 | CLSB |  | 1N044, IN046, IN077 |  | Pipette, Pipettman Neo P1000N |  | Gilson | F144565 | [44 empty] | 36 |  |  | Group III | Lab
|  | PSU Micro/Cell/Recomb DNA Labs | 2323 | CLSB |  | 1N044, IN046, IN078 |  | Pipette, Pipettman Neo P10N |  | Gilson | F144565 | [44 empty] | 36 |  |  | Group III | Lab
|  | PSU Micro/Cell/Recomb DNA Labs | 2324 | CLSB |  | 1N044, IN046, IN079 |  | Pipette, Pipettman Neo P2N |  | Gilson | F144565 | [44 empty] | 36 |  |  | Group III | Lab
|  | PSU | 2325 | CLSB | 3 | TBD | TBD | Phosphorimager - Typhoon FLA 9500 and analysis software |  | GE Healthcare Life Sciences | 29-0040-80 Typhoon FLA 9500 | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU | 2326 | CLSB |  | TBD | TBD | Floor Centrifuges |  | Thermo Fisher Sorvall | RC3BP | [44 empty] | 1 |  |  | Group III | Lab
|  | Dr. Gray | 2327 |  | P2 |  |  | UPS GE LP8-11u with 2 extra battery cabinets for 1 hr battery life |  | General Electric | TBD | [44 empty] | 1 |  |  | Group III | Computer
|  | Dr. Gray | 2328 |  | P2 | P2N021 |  | CryoPlus2 model 7402 |  | Thermo Scientific | model 7402 | [44 empty] | 1 |  |  | Group III | Lab
|  | Dr. Gray | 2329 |  | P2 | P2N021 |  | Freezer, -80C, Ultra Low 28.8 Cu. Ft. No Alarm |  | Thermo Scientific | UxF60086A, 120V | [44 empty] | 1 |  |  | Group III | Lab
|  | Dr. Gray | 2341 |  | P2 | P2N021DA |  | Labconco Biosafety cabinet, 6' #3460009 |  | Labconco | Labconco Purifier Delta Series # 3460009 6' with 7370600 base | [44 empty] | 2 |  |  | Group III | Lab
|  | Dr. Gray | 2342 |  |  | P2N021 |  | General low speed lab shaker |  | ORBIT / Spectra Services | TBD |  |  | 12.6 | 6.3 | 10 |  |  |  |  | 120 | 60 | [33 empty] | 1 |  |  | Group III | Lab
|  | Dr. Gray | 2343 |  | P2 | P2N021DA |  | Inverted Microscope |  | Zeiss or Nikon | Zeiss PrimoVert or Nikon Eclipse TS100F | [44 empty] | 1 |  |  | Group III | Lab
|  | PSU | 2344 | CLSB |  | TBD | TBD | Freezer, -80C, with no alarm system s |  | Thermo Scientific GP series | TBD | Floor | 35 | 32 | 79 | Min. | Open | Min. | Open | 429 | 120 | 10? | 1 | Y | N | Y | N | Y | 5-15P | N | Y | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | [5 empty] | 1 |  |  | Group III | Lab